®

CUSTOMER SUPPORT SERVICES

Fluke-741B, 753B,744 & Honeywell 2020 Parts Listings Reference Designations Schematics

Revision

6-Nov-08

Company Proprietary Information

Fluke Corporation

Everett, Washington

This is the Service Documentation for all Fluke-744, for 741B and 743B with SNs above 7932001 (V2.0 or greater firmware) and 2020 Documenting Process Calibrator (with V4.0 or higher firmware).  There have been 2 generations of 744 Main PCAs, the first used a Fluke custom chip, the second uses standard parts. This package includes Parts Listings for all boards, Reference Designation drawings, and both generations of Schematics.

These drawings are Company Proprietary Information.  They are not intended for customer distribution.

Please safeguard these drawings.  A 741B, 743B, 744 or 2020 Service Manual is not available.  There are no plans to produce a 741B, 743B, 744 or 2020 Service Manual.

The performance verification procedure and calibration procedure for the 741B, 743B, 744 or 2020 are included in the Fluke-7xx Calibration Manual, it is available under P/N 602505.

The 2020 Hart is a special version of the 744 made for Honeywell.  It has a different case and different firmware.   The PCAs are the same as the 744.  For 2020's with Firmware versions of less than V3.3 should be serviced using the 743 information from SA 1501.

Final Assembly Fluke-741B, 743B, 744, 2020

| Item Seq | 741B PN | 743B PN | 744 PN | 2020 PN | Description | Quantity |
| --- | --- | --- | --- | --- | --- | --- |
| 2 | 1607554 | 1607568 | 662783 | 662783 | Tested Board Set, Main, PS and I/O PCAs | 1 |
| 6 | 202421 | 202421 | 202421 |  | Top Case, Grey/Yellow | 1 |
| 6 |  |  |  | 938345 | Top Case, Grey | 1 |
| 7 | 202439 | 202439 | 202439 |  | Bottom Case, Grey/Yellow | 1 |
| 7 |  |  |  | 202488 | Bottom Case, Grey | 1 |
| 8 | 946681 | 946681 | 946681 | 946681 | Display Bracket | 1 |
| 10 | 938357 | 938357 | 938357 | 938357 | Battery Door | 1 |
| 12 | 938274 | 938274 | 938274 | 938274 | Side Plug | 1 |
| 13 | 3369101 | 3369101 | 3369101 | 3369101 | Mylar Insulator | 1 |
| 24 | 949487 |  |  |  | Lens With Logo, 741B | 1 |
| 24 |  | 949490 |  |  | Lens With Logo, 743B |  |
| 24 |  |  | 662871 |  | Lens With Logo, 744 | 1 |
| 24 |  |  |  | 202405 | Lens With Logo, 2020 | 1 |
| 26 |  |  |  | 202389 | PCA, Keyboard, A36, CE Vers | 1 |
| 27 | 884254 | 884254 | 884254 | 884254 | Insulator, Input Receptacle (Red) | 3 |
| 28 | 884259 | 884259 | 884259 | 884259 | Insulator, Input Receptacle (Black) | 3 |
| 50 | 202330 | 202330 | 202330 | 202330 | Keyboard Shield | 1 |
| 51 | 938170 | 938170 |  |  | Battery Pack Assy, NiCad |  |
| 51 |  |  | 665083 | 665083 | Battery Pack, NiMh | 1 |
| 52 |  |  | 662863 | 662863 | Keypad, 744 & 2020 | 1 |
| 52 | 606910 | 606910 |  |  | Keypad 741B & 743B | 1 |
| 55 | 202413 | 202413 | 202413 |  | Top Case Decal, Yellow | 1 |
| 55 |  |  |  | 946756 | Top Case Decal, 2020 | 1 |
| 58 | 949479 | 949479 | 949479 | 949479 | Label, Certified Calibration | 1 |
| 100 | 3378798 | 3378798 | 3378798 | 3378798 | C-Tick/Weee Sticker, Agency And Weee Decal | 1 |
| 821 | 3323456 | 3323456 | 3323456 | 3323456 | Main Shield | 1 |
| 823 | 929369 | 929369 | 929369 | 929369 | Battery,Primary,Lithium-Mno2, 3v, 0.18ah, CR2032, Coin | 1 |
| 824 | 602430 | 602430 | 602430 | 602430 | LCD Module, 240x200, Monochrome Graphics, Led Backlight, 0.29mm Dpi | 1 |
| 825 | 601051 | 601051 | 601051 | 601051 | Ground Strip, BeCu, Spring Fingers,.76 | 1 |
| 833 | 494641 | 494641 | 494641 | 494641 | Screw, 5-20,.312, Washer Head, Phillips, Steel, Zinc-Chromate | 10 |
| 834 | 942797 | 942797 | 942797 | 942797 | Screw, Btnh,T,SS,M3.5x16 | 6 |
| 836 | 256164 | 256164 | 256164 | 256164 | Screw, Ph, P, Lock, Ss, 4-40, .375 | 2 |
| 838 | 601390 | 601390 | 601390 | 601390 | Washer, Flat, Silicone, .200, .500, .032 | 6 |
| 844 | 2726702 | 2726702 | 2726702 | 2726702 | Cable Assembly, Flat, 2 Row, 26 Cond,1.25 L | 1 |
| 857 | 949441 | 949441 | 949441 | 949441 | Battery Pad | 1 |
| 858 | 105841 | 105841 | 105841 | 105841 | Tape, Foam, Urethane, Adhes, .063, .250 | 0.25 |

Accessories

| Item Number | Description | Quantity |
| --- | --- | --- |
| 927579 | Safety Alligator, Black, AC20 | 2 |
| 927777 | Banana Test Probe, Red,TP20 | 1 |
| 689653 | Cable Assembly, 2 Cond,9p/9s D-Sub, | 1 |
| 927772 | Banana Test Probe, Black,TP20 | 1 |
| 927798 | Test Lead, Red, R/A Plug - Plug | 2 |
| 927793 | Test Lead, Black, R/A Plug - Plug | 2 |
| 944632 | Test Lead, 10cm, Black | 2 |
| 946769 | Strap, 701/702 | 1 |
| 604824 | Information Pack, Manual/Disk, Dpc/Track Demo | 1 |
| 943738 | Cable Assembly, 7 Cond, 9s Dsub, 9p Dsub | 1 |
| 691862 | Instruction Sheet Set, BC7217 | 1 |
| 691854 | Battery Charger, Ni-Cd/Mh, 7.2V, 1.7A | 1 |
| 690344 | Power Supply, 23W,15V@1.5A,Charger/Converter, USA | 1 |
| 690552 | Power Supply, 23W,15V@1.5A,Charger/Converter, International | 1 |
| 603016 | 2020 Users Manual |  |
| 802360 | 2020 Hart Mode Users Guide |  |
| 804341 | 744 Users Manual Set |  |
| 649787 | 741B & 743B Users Manual Set |  |

First Version Main PCA Parts

| Ref Des | 741B PN | 743B PN | 744, 2020 PN | Description |
| --- | --- | --- | --- | --- |
| BP1-6 | 665130 | 665130 | 665130 | Terminal Post, 74x |
| BT1 | 929448 | 929448 | 929448 | Holder, Battery, Pwb, 2032 Coin |
| BT1 | 929369 | 929369 | 929369 | Battery,Primary,Lithium-MnO2, 3v, 0.18ah, CR2032, Coin |
| C1 | 876552 | 876552 | 876552 | Capacitor, Electrolytic, Tantalum, 3.3uF, +-20%, 16v, 3528 |
| C2 C48 C103 C184-186 C307 | 740597 | 740597 | 740597 | Capacitor,Ceramic,0.22uF,+80-20%,50v,Y5V,1206 |
| C3 | 867580 | 867580 | 867580 | Capacitor, Tant, 47uF, +-20%, 10v, 7343 |
| C4 | 855101 | 855101 | 855101 | Capacitor, Ceramic, 22pF,+-5%,50v,COG,0805 |
| C5 C7 C11 C17 C20 C25 C29 C37 C212-213 | 867650 | 867650 | 867650 | Capacitor, Cer,100pF,+-1%,50v,COG,0805 |
| C6 C28 C45-46 C104 | 691508 | 691508 | 691508 | Capacitor,Ceramic,100pF,+-10%,500v,COG,1206 |
| C8 | 758078 | 758078 | 758078 | Capacitor, Cer,220pF,+-10%,50v,COG,1206 |
| C9 C36 | 927442 | 927442 | 927442 | Capacitor, Cer,6.8pF,+-0.5pF,50v,COG,0805 |
| C10 C15 C30 | 106146 | 106146 | 106146 | Capacitor,Ceramic,0.01uF,+-10%,50v,X7R,0805 |
| C12 C57 | 769240 | 769240 | 769240 | Capacitor, Cer,33pF,+-10%,50v,COG,1206 |
| C122-123 | 867093 | 867093 | 867093 | Capacitor, Polypropylene, 0.047uF, 10%, 63v, 5mm Radial |
| C124 | 887278 | 887278 | 887278 | Capacitor,Ceramic,390pF,+-5%,50v,COG,1206 |
| C125 | 929302 | 929302 | 929302 | Capacitor, Smr,Cap,Ta,1.5uF,+-20%,50v,6032 |
| C126 | 721050 | 721050 | 721050 | Capacitor, Polypropylene, 0.033uF, +-10%, 63v, 5mm Radial |
| C127-128 | 514174 | 514174 | 514174 | Capacitor,Ceramic,15pF,+-5%,50v,COG,0805 |
| C13-14 C35 C129-130 C133-134 C315 C317 | 942529 | 942529 | 942529 | Capacitor,Ceramic,0.1uF,+-10%,25v,X7R,0805 |
| C135 C214 | 782615 | 782615 | 782615 | Capacitor, Cer,0.047uF,+-20%,50v,X7R,1206 |
| C16 C51-52 C120 | 927814 | 927814 | 927814 | Capacitor,ElectrolytIC,Tantalum,10uF,+-20%,25v,6032 |
| C18 C116 C302 | 484378 | 484378 | 484378 | Capacitor, Cer,1000pF,+-10%,50v,X7R,0805 |
| C19 C23 C58 C60 C113 C201 | 769240 | 769240 | 769240 | Capacitor, Cer,33pF,+-10%,50v,COG,1206 |
| C21-22 C136-139 C144 C147-148 C150 C153-156 C158 C160 C162-168 C171 C174-175 C194-197 C223-224 | 106625 | 106625 | 106625 | Capacitor, Cer,0.22uF,+-10%,25v,X7R,1206 |
| C24 | 867650 | 867650 | 867650 | Capacitor, Cer,100pF,+-1%,50v,COG,0805 |
| C26 C118-119 | 876545 | 876545 | 876545 | Capacitor,Tantalum,22uF,+-20%,6v,6032 |
| C27 C41 C49-50 C117 C142 C222 C308 C310-311 C314 | 740597 | 740597 | 740597 | Capacitor,Ceramic,0.22uF,+80-20%,50v,Y5v,1206 |
| C31-32 C140-141 C143 C145-146 C151-152 C161 C169 C172-173 C193 C199 | 106625 | 106625 | 106625 | Capacitor, Cer,0.22uF,+-10%,25v,X7R,1206 |
| C33 | 832279 | 832279 | 832279 | Capacitor, Cer,4700pF,+-10%,50v,X7R,1206 |
| C34 C43 C47 | 930235 | 930235 | 930235 | Capacitor, Cer,47pF,+-10%,1000v,COG,1808 |
| C38 | 691516 | 691516 | 691516 | Capacitor, Cer,0.01uF,+-10%,500v,X7R,1210 |
| C39-40 | 930235 | 930235 | 930235 | Capacitor, Cer,47pF,+-10%,1000v,COG,1808 |
| C42 C106 C318 | 106146 | 106146 | 106146 | Capacitor, Ceramic,0.01uF,+-10%,50v,X7R,0805 |
| C44 C207 | 927442 | 927442 | 927442 | Capacitor, Cer,6.8pF,+-0.5pF,50v,COG,0805 |
| C53-56 C105 C220-221 | 691508 | 691508 | 691508 | Capacitor, Ceramic,100pF,+-10%,500v,COG,1206 |
| C59 C202-203 C309 | 942529 | 942529 | 942529 | Capacitor, Ceramic,0.1uF,+-10%,25v,X7R,0805 |
| C101-102 C182 C191-192 | 855101 | 855101 | 855101 | Capacitor, Ceramic, 22pF,+-5%,50v,COG,0805 |
| C300 |  |  | 484378 | Capacitor, Cer, 1000pF,+-10%,50v,X7R,0805 |
| C301 |  |  | 943365 | Capacitor, Ceramic, 470pF,+-5%,50v,COG,1206 |
| C304 |  |  | 875914 | Capacitor, Cer,200pF,+-5%,50v,COG,0805 |
| C319 | 2131344 | 2131344 | 2131344 | Capacitor, Ceramic, 10 UF,+-10%,6.3v,X5R,0805 |
| CR2 CR102 CR105 CR107 CR110 CR300-301 | 742320 | 742320 | 742320 | Diode,Si,Pn,BAV99,70v,215ma,6ns,Dual,Series,SOT-23 |
| CR100 | 830489 | 830489 | 830489 | Diode,Si,BAS16, 75v,200ma,6ns,SOT-23 |
| CR103-104 | 107573 | 943642 | 107573 | Diode, Si, Pn, 400v,1a,1.8us,Do-214AC |
| CR108-109 | 929745 | 929745 | 929745 | Diode,Si,Schottky,Bat54s,30v,200ma,5ns,Dual-Series,SOT-23 |
| J1-4 | 943709 | 943709 | 943709 | Header, 2 Row, 2mm Ctr,20 Pin |
| J5 | 929229 | 929229 | 929229 | Connector, Flat Flex, 1mm Ctr, Rt Ang,16 Pos |
| J7 |  |  |  | Fluke-744-4035 PCA, Thermal |
| K1-2 K4 | 942917 | 942917 | 942917 | Relay, Armature, 2 Form C, 5Vdc,Latch |
| L2 L7 L14 | 601010 | 601010 | 601010 | Inductor, Bead, 600 Ohms@100mhz, 200mAdc, 1ohm, 1206 |
| L3-6 L8 L13 L15-16 | 601010 | 601010 | 601010 | Inductor, Bead, 600 Ohms@100mhz, 200madc, 1ohm, 1206 |
| LS1 | 876995 | 876995 | 876995 | Audio Transducer, Piezo, Sounder, 2kHz, 10V, Washable, Trimmed Leads, 26x22mm |
| MP1 MP2 | 745802 | 745802 | 745802 | Fluke 52-8004, Contact, Thermocouple |
| Q1 Q5 Q106 | 742676 | 742676 | 742676 | Transistor, Npn, MMBT3904, 60V, 350mW, SOT-23 |
| Q2 Q100 Q107 | 742684 | 742684 | 742684 | Transistor, Si, Pnp, Mmbt3906, 40v, 200mA, 250mHz, 225mW, SOT-23 |
| Q3-4 | 876263 | 876263 | 876263 | Transistor, Si, N-Jfet, Sel, SOT-23 |
| Q105 | 782490 | 782490 | 782490 | Transistor, N-Ch R05A, Si, N-Dmos, 500v, TO-92 |
| Q108 | 782508 | 782508 | 782508 | Mosfet, P-Chn R05A, Transistor,Si,P-Mos,500v,TO-92 |
| Q200-201 |  | 107573 | 643681 | Transistor, Si, N-Mos, NDS351AN, 500mW, SOT-23 |
| R1-2 R5-6 R29 R40 R62 R107-108 R113 R115 R220 R308 | 928791 | 928791 | 928791 | Resistor,Cermet,10k,+-1%,0.125W, 100ppm, 0805 |
| R3 | 928825 | 928825 | 928825 | Resistor,Cermet,24.9k,+-1%,0.125W, 100ppm, 0805 |
| R4 R10 R26 R50 R58 R168 R319 | 928713 | 928713 | 928713 | Resistor,Cermet,1k,+-1%,0.1W, 100ppm, 0805 |
| R7 R9 | 944236 | 944236 | 944236 | Resistor,Mf,100k,+-0.1%,0.125W, 25ppm, 1206 |
| R8 R18 R32 R52 R73-74 R122-127 R315 | 928866 | 928866 | 928866 | Resistor,Cermet,100k,+-1%,0.125W, 100ppm, 0805 |
| R11 R156 R162 R236 | 928762 | 928762 | 928762 | Resistor, Cerm,4.02k,+-1%,0.1W, 100ppm, 0805 |
| R12 | 928924 | 928924 | 928924 | Resistor,Cermet,10,+-1%,0.125W, 100ppm, 0805 |
| R13 | 831305 | 831305 | 831305 | Resistor, Cerm,56.2k,+-1%,.125W, 100ppm, 1206 |
| R14 R106 R109-111 R121 R161 R223 | 807172 | 807172 | 807172 | Resistor, Cerm,2k,+-1%,.125W, 100ppm, 1206 |
| R15 R312 R321 R49 R118 | 928866 | 928866 | 928866 | Resistor,Cermet,100k,+-1%,0.125W, 100ppm, 0805 |
| R16-17 R28 R33 R57 R66 R68-69 R72 R80 R116 R120 R129 R147 R150 R167 R219 R233 R239 R310-311 | 928791 | 928791 | 928791 | Resistor,Cermet,10k,+-1%,0.125W, 100ppm, 0805 |
| R19 | 867085 | 867085 | 867085 | Resistor, Cerm,75k,+-1%,.125W, 100ppm, 1206 |
| R20-21 | 928908 | 928908 | 928908 | Resistor, Cerm,66.5k,+-1%,0.1W, 100ppm, 0805 |
| R23 | 928820 | 928820 | 928820 | Resistor,Cermet,20k,+-1%,0.125W, 100ppm, 0805 |
| R24 R45 | 783274 | 783274 | 783274 | Resistor,Cermet,10m,+-5%,0.25W, 400ppm, 1206 |
| R31 | 928937 | 928937 | 928937 | Resistor,Cermet,100,+-1%,0.125W, 100ppm, 0805 |
| R34 | 867692 | 867692 | 867692 | Resistor, Cerm,40.2k,+-1%,.125W, 100ppm, 1206 |
| R35 R171 R323 | 867432 | 867432 | 867432 | Resistor, Cerm,562k,+-1%,.125W, 100ppm, 1206 |
| R36 R306 | 944285 | 944285 | 944285 | Resistor, Cerm,499k,+-1%,0.1W, 100ppm, 0805 |
| R37 R41 | 928945 |  | 928945 | Resistor,Cermet,1m,+-1%,0.125W, 100ppm, 0805 |
| R38 | 943311 | 943311 | 943311 | Resistor,Cermet,44.2k,+-1%,0.125W, 100ppm, 0805 |
| R42 R54 R119 | 928882 | 928882 | 928882 | Resistor,Cermet,200k,+-1%,0.1W, 100ppm, 0805 |
| R43 | 691920 | 691920 | 691920 | Resistor, MF, 39.2k,+-1%,0.1W, 50ppm, 0805 |
| R44 R112 | 928713 | 928713 | 928713 | Resistor,Cermet,1k,+-1%,0.1W, 100ppm, 0805 |
| R46 R55 R169 | 867460 | 867460 | 867460 | Resistor, Cerm,665k,+-1%,.125W, 100ppm, 1206 |
| R47 | 687530 | 687530 | 687530 | Resistor, Metal Film,4.99k,\+-0.5%,0.125W, 10ppm, 1206 |
| R48 R155 | 928825 | 928825 | 928825 | Resistor,Cermet,24.9k,+-1%,0.125W, 100ppm, 0805 |
| R51 | 811968 | 811968 | 811968 | Resistor,Cermet,15m,+-5%,0.25W, 300ppm, 1206 |
| R53 R75 R79 | 867692 | 867692 | 867692 | Resistor, Cerm,40.2k,+-1%,.125W, 100ppm, 1206 |
| R59 | 930110 | 930110 | 930110 | Resistor,Cermet,5.62k,+-1%,0.125W, 100ppm, 0805 |
| R60 | 928861 | 928861 | 928861 | Resistor, Cerm,61.9k,+-1%,0.1W, 100ppm, 0805 |
| R61 R82 | 2100621 | 2100621 | 2100621 | Resistor,Variable,Cermet,50k,+25,-0%, 0.2W, 250ppm, 4mm Smt, Top Adj |
| R63 | 810630 | 807172 | 810630 | Resistor, Cerm,1.5k,+-1%,.125W, 100ppm, 1206 |
| R67 | 928820 | 928820 | 928820 | Resistor,Cermet,20k,+-1%,0.125W, 100ppm, 0805 |
| R71 | 807735 | 807735 | 807735 | Resistor, Cerm,27,+-5%,.125W, 200ppm, 1206 |
| R76 | 877006 | 877006 | 877006 | Resistor, Cerm,402k,+-1%,.125W, 100ppm, 1206 |
| R77 | 944835 | 944835 | 944835 | Resistor, Var,Cerm,10k,+-25% |
| R78 | 801258 | 801258 | 801258 | Resistor, Cerm,30.1k,+-1%,.125W, 100ppm, 1206 |
| R81 R234 | 928994 | 928994 | 928994 | Resistor,Cermet,4.75m,+-1%,0.125W, 400ppm, 0805 |
| R83 | 822312 | 822312 | 822312 | Resistor, Cerm,1.87k,+-1%,.125W, 100ppm, 1206 |
| R101 R157-158 | 810630 | 810630 | 810630 | Resistor, Cerm,1.5k,+-1%,.125W, 100ppm, 1206 |
| R102 R114 R117 | 928945 | 928945 | 928945 | Resistor,Cermet,1m,+-1%,0.125W, 100ppm, 0805 |
| R103-104 R235 | 867697 | 867697 | 867697 | Resistor, Cerm,150k,+-1%,.125W, 100ppm, 1206 |
| R105 | 807172 | 928762 | 807172 | Resistor, Cerm,2k,+-1%,.125W, 100ppm, 1206 |
| R128 | 943642 | 810630 | 943642 | Resistor, Cerm,475,+-1%,0.1W, 100ppm, 0805 |
| R142-143 | 746644 | 746644 | 746644 | Resistor, Cerm,20k,+-5%,.125W, 200ppm, 1206 |
| R145-146 R152 | 876524 | 876524 | 876524 | Resistor, Cerm,226k,+-1%,.125W, 100ppm, 1206 |
| R154 | 688838 | 688838 | 688838 | Resistor,Bulk Metal Foil,200,+-1%,0.25W, 3ppm,6032 |
| R159-160 | 811885 | 811885 | 811885 | Resistor,Cermet,432,+-1%,0.25W, 100ppm, 1206 |
| R166 | 821181 | 821181 | 821181 | Resistor,Cermet,1.5m,+-1%,0.125W, 200ppm, 1206 |
| R170 | 886833 | 886833 | 886833 | Resistor,Cermet,280k,+-1%,0.25W, 100ppm, 1206 |
| R172 | 928861 | 928861 | 928861 | Resistor, Cerm,61.9k,+-1%,0.1W, 100ppm, 0805 |
| R173 | 928817 | 928817 | 928817 | Resistor, Cerm,16.9k,+-1%,0.1W, 100ppm.0805 |
| R174 | 929039 | 929039 | 929039 | Resistor, Cerm,845,+-1%,0.1W, 100ppm, 0805 |
| R180 R226-232 R313 | 927707 | 927707 | 927707 | Resistor, Cerm,47,+-5%,.0625W, 200ppm,0603 |
| R224-225 | 928762 | 928945 | 928762 | Resistor, Cerm,4.02k,+-1%,0.1W, 100ppm, 0805 |
| R237 | 944236 | 944236 | 944236 | Resistor,Mf,100k,+-0.1%,0.125W, 25ppm, 1206 |
| R238 | 746347 | 746347 | 746347 | Resistor, Cerm,220,+-5%,.125W, 200ppm, 1206 |
| R300 |  |  | 877006 | Resistor, Cerm,402k,+-1%,.125W, 100ppm, 1206 |
| R301 R305 |  |  | 866830 | Resistor, Cerm,115k,+-1%,.125W, 100ppm, 1206 |
| R302 |  |  | 887112 | Resistor, Cerm,487k,+-1%,.125W, 100ppm, 1206 |
| R303 |  |  | 821678 | Resistor,Cermet,499k,+-1%,0.25W, 100ppm, 1206 |
| R304 |  |  | 1549425 | Resistor,Cermet,215k,+-1%,0.1W, 100ppm, 0805 |
| R307 |  |  | 836437 | Resistor, Cerm,3M,+-5%,.125W, 200ppm, 1206 |
| R309 |  |  | 928924 | Resistor,Cermet,10,+-1%,0.125W, 100ppm, 0805 |
| R314 |  |  | 836643 | Resistor, Cerm,215k,+-1%,.125W, 100ppm, 1206 |
| R316 |  |  | 927421 | Resistor,Cermet,20k,+-1%,0.25W, 100ppm, 1206 |
| R317 |  |  | 769802 | Resistor, Cerm,100k,+-1%,.125W, 100ppm, 1206 |
| R318 |  |  | 746768 | Resistor, Cerm,300k,+-5%,.125W, 200ppm, 1206 |
| R320 |  |  | 816033 | Resistor, Cerm,14k,+-1%,.125W, 100ppm, 1206 |
| R322 |  |  | 944439 | Resistor,Cermet,0.22,+-5%,0.125W, 600ppm, 0805 |
| U1 | 867932 | 867932 | 867932 | IC, Op Amp, Dual, LM358, Low Power, Soic |
| U2 | 929315 | 929315 | 929315 | IC, Op Amp, TC913, 30uV Offset, Dual, Chopper-Stab, SO8 |
| U3 | 913835 | 913835 | 913835 | IC, Cmos, Comparator TLC393, Dual, uPower, SO8 |
| U4 | 929828 | 929828 | 929828 | IC, Op Amp, AD820A, 5V To 36V, 0.8mV Offset, 1.8mHz, R/R, S/S, Lo Pwr, SO8 |
| U5 U10-12 | 876503 | 876503 | 876503 | IC, Cmos LTC1050 Op Amp, Chopper, Int Caps, SO8 |
| U6 | 2570483 | 2570483 | 2570483 | IC, Programmable Waveform Generator,AD9833, 2.3-5.5V, Msop10 |
| U7 | 107565 | 107565 | 107565 | IC, Cmos, LTC1150 Op Amp, Chopper, +/-15v, Int Caps, SO8 |
| U8-9 U101-102 | 914101 | 914101 | 914101 | IC, Memory, Sram, K6T1008, 1MB, 128kx8, 5v, 100ns, SO32 |
| U13-16 | 876503 | 876503 | 876503 | IC, Cmos LTC1050 Op Amp, Chopper, Int Caps, SO8 |
| U100 | 929083 | 929083 | 929083 | IC, Microprocessor, 68306, 16 Bit, 5v, 16mHz, Integrated, PQFP132 |
| U103-104 | 107562 | 107562 | 688827 | IC, Memory, Flash, 28F008BV-B, 8MB, 1Mx8, 5V, 80ns, Bottom Boot, TSOP40 |
| U106 | 914036 | 914036 | 914036 | Real Time Clock Module, RTC64613A, Parallel I/O, W/Crystal, So24 |
| U107 | 929208 | 929208 | 929208 | IC, LCD Driver, S1D13305, Graphic LCD Controller, 3v, QFP60 |
| U108 | 929609 | 929609 | 929609 | IC, Memory, Sram, 43256, 256kb, 32kx8, 5v, 70ns, SO28 |
| U109-110 | 852090 | 852090 | 852090 | IC, Cmos 74HC126 Quad Buffer W/3-St, Soic |
| U111-112 | 852090 | 852090 | 852090 | IC, Cmos 74HC126 Quad Buffer W/3-St, Soic |
| U113 | 866801 | 866801 | 866801 | IC, Cmos 74HC125 Quad Bus Buffer W/3-St, Soic |
| U114 | 929224 | 929224 | 929224 | IC, Cmos 694 Microprocessor Supervisor, SO8 |
| U115 | 866785 | 866785 | 866785 | IC, Cmos, 68C681,Dual Channel Uart, Plcc |
| U117 | 867932 | 867932 | 867932 | IC, Op Amp LM358 Dual, Low Power, Soic |
| U118 | 807040 | 807040 | 807040 | IC, Bipolar 1403 2.5v,40 Ppm T.C.,Bandgap Ref, Soic |
| U119 | 929273 | 929273 | 929273 | IC Cmos 74HC4094 8bit Shft Rgs W/3-St, Soic |
| U120 | 782995 | 782995 | 782995 | IC, Cmos, 74HC74 Dual D F/F,+Edg Trg, Soic |
| U121 | 929299 | 929299 | 929299 | IC, Cmos, MAX529 8 Bit Ser Dac,1 LSB, Soic |
| U123 | 876813 | 876813 | 876813 | IC, Cmos 74ACT14 Hex Inverter W/Scht Trig, Soic |
| U124 | 929075 | 929075 | 929075 | IC, Op Amp 34072 Dual, High Bw, Sngl Sup, SO8 |
| U125 | 642436 | 642436 | 642436 | IC, Analog Switch, ADG433, 12 To -15V, 24 Ohms, Spst, Quad, No, Isolated, So16 |
| U126 | 107524 | 107524 | 107524 | IC, Cmos ADG513 Quad Bilateral Switch, Soic |
| U128 | 2132280 | 2132280 | 2132280 | IC, Dac, LTC1590, 5.0v,12-Bit, Dual, Current, Serial, Soic16 |
| U129 U170 | 806794 | 806794 | 806794 | IC,Op Amp TI062 Dual, Jfet, Low Power, Soic |
| U130 U134 | 107524 | 107524 | 107524 | IC, Cmos ADG513 Quad Bilateral Switch,Soic |
| U135 | 942701 | 942701 | 942701 | IC, Cmos, ADG511 Quad Bilateral Switch,Soic |
| U136 | 806794 | 806794 | 806794 | IC,Op Amp TI062 Dual, Jfet, Low Power, Soic |
| U137 | 868000 | 868000 | 868000 | IC, Bipolar 737J True Rms To Dc Converter, Soic |
| U140 | 776195 | 776195 | 776195 | IC, Meas, Merc5 Processor & A/D Conv, Cmos |
| U141 U144 | 942680 | 942680 | 942680 | Isolator, Opto, HCPL:-0501, LED To Transistor, Single Channel,SO8 |
| U142-143 U162 | 929286 | 929286 | 929286 | Isolator, Opto, LED To Darlington, Soic |
| U158 | 838193 | 838193 | 838193 | IC,Cmos 74AC02 Quad 2 Input Nor Gate, Soic |
| U163 | 942698 | 942698 | 942698 | IC,Voltage, Bipolar 1021 Ref,7 V,+-50 Mv,10 Ppm, SO8 |
| U164 | 942698 | 942698 | 942698 | IC,Voltage, Bipolar 1021 Ref,7 V,+-50 Mv,10 Ppm, SO8 |
| U165 | 929278 | 929278 | 929278 | IC,Cmos 74HC4051 8-1 Line Mux Analog Sw, Soic |
| U166 | 929273 | 929273 | 929273 | IC, Cmos, 74HC4094 8bit Shft Rgs W/3-St, Soic |
| U171 | 837054 | 837054 | 837054 | IC, Logic, 74HC393, 2-6V, Divide By 16 Binary Counter, Dual, SO14 |
| U172 |  |  | 2061960 | IC, Comm, A5191HRT, Hart Modem, 3.3V-5.0V, Plcc28 |
| U173 U175 | 866801 | 866801 | 866801 | IC,Cmos, 74HC125 Quad Bus Buffer W/3-St, Soic |
| U174 |  |  | 929489 | IC,Voltage Reference, LM4041, 1.2V, 2%, 150ppm/C, 0.065-20mA, SOT-23 |
| VR104-105 | 837161 | 837161 | 837161 | Zener, Uncomp, MMBZ5233B, 6V, 5%, 20mA, 225mA, SOT-23 |
| VR3-4 | 837187 | 837187 | 837187 | Zener, Uncomp, MMBZ5245B, 15W, 5%, 8.5mA, 225mW, SOT-23 |
| WP1-2 | 2563832 | 2563832 | 2563832 | Standoff, 4-40, .187, Steel, Tin, Smd |
| Y1 | 929245 | 929245 | 929245 | Crystal, 3.84mHz, 50/100ppm, 20pF, Plastic Encapsulated, Smd |
| Y2 | 929240 | 929240 | 929240 | Crystal, 3.6864mHz, 100/100ppm, 20pF, Plastic Encapsulated, Smd |
| Y3 | 943105 | 943105 | 943105 | Crystal, 10mhz,50/100ppm,20pF,Plastic Encapsulated |
| Z3 | 833921 | 833921 | 833921 | Rnet, Mf, Poly, Sip, Fluke 45 A To D Conv |
| Z4-5 | 900451 | 900451 | 900451 | Fluke-700-4r04t-P ,Rnet, Mf, Poly, Dip, Hi V Inst Amp |
| Z6 | 103143 | 103143 | 103143 | Fluke-700-4r05t-P ,Rnet, Mf, Poly, Dip, Lo V Inst Amp |
| Z11-12 | 626462 | 626462 | 626462 | Fluke-744-8101 Smr,Res Net, Thin Film, Tested, Smr |

Second Version Main PCA Assembly

| Ref Des | 741B PN | 743B PN | 744 2020 PN | Description |
| --- | --- | --- | --- | --- |
| BP1-6 | 665130 | 665130 | 665130 | Terminal Post, 74x |
| C1 | 876552 | 876552 | 876552 | Capacitor, Tantalum, 3.3uF, +-20%, 16v, 3528 |
| C3 | 867580 | 867580 | 867580 | Capacitor, Tant, 47uF, +-20%,10v, 7343 |
| C5 C17 | 605359 | 605359 | 605359 | Capacitor, Cer,100PF,+-10%,50V,C0G,0603 |
| C6 C28 C53-56 C105 C220 C221 | 691508 | 691508 | 691508 | Capacitor,Ceramic,100pF,+-10%,500v,COG,1206 |
| C7 C25 C29 C212 C213 | 867650 | 867650 | 867650 | Capacitor, Cer,100pF,+-1%,50v,COG,0805 |
| C8 | 758078 | 758078 | 758078 | Capacitor, Cer,220pF,+-10%,50v,COG,1206 |
| C9 | 927442 | 927442 | 927442 | Capacitor, Cer,6.8pF,+-0.5pF,50v,COG,0805 |
| C10 C15 C316 C318 | 106146 | 106146 | 106146 | Capacitor,Ceramic,0.01uF,+-10%,50v,X7R,0805 |
| C11 | 605359 | 605359 | 605359 | Capacitor, Cer,100PF,+-10%,50V,C0G,0603 |
| C12 C57 | 769240 | 769240 | 769240 | Capacitor, Cer,33pF,+-10%,50v,COG,1206 |
| C13 C14 | 942529 | 942529 | 942529 | Capacitor,Ceramic,0.1uF,+-10%,25v,X7R,0805 |
| C16 C52 | 927814 | 927814 | 927814 | Capacitor,ElectrolytIC,Tantalum,10uF,+-20%,25v,6032 |
| C18 C116 C302 |  |  | 484378 | Capacitor, Cer,1000pF,+-10%,50v,X7R,0805 |
| C19 C23 C58 C60 C113 C201 | 769240 | 769240 | 769240 | Capacitor, Cer, 33pF,+-10%,50v,COG,1206 |
| C20 | 1579884 | 1579884 | 1579884 | Capacitor, Cer, 100PF,+-5%,100V,C0G,0603 |
| C21 C22 C136-139 C144 C147 C148 C150 C153-156 C158 C160 C162-168 C171 C194-197 | 106625 | 106625 | 106625 | Capacitor, Cer,0.22uF,+-10%,25v,X7R,1206 |
| C24 | 867650 | 867650 | 867650 | Capacitor, Cer,100pF,+-1%,50v,COG,0805 |
| C26 | 876545 | 876545 | 876545 | Capacitor,Tantalum,22uF,+-20%,6v,6032 |
| C27 C72 C82 C83 | 2151309 | 2151309 | 2151309 | Capacitor, Cer, 1UF,+-10%,10V,X5R,0603 |
| C30 C42 | 1612566 | 1612566 | 1612566 | Capacitor, Film, 0.1UF,+-10%,50V,5750,TAPE |
| C31 C32 C140-143 C145 C146 C151 C152 C161 C169 C172 | 106625 | 106625 | 106625 | Capacitor, Cer,0.22uF,+-10%,25v,X7R,1206 |
| C33 | 2091994 | 2091994 | 2091994 | Capacitor, Cer, 4700PF,+-10%,50V,X7R,0603 |
| C34 C43 C47 | 930235 | 930235 | 930235 | Capacitor, Cer,47pF,+-10%,1000v,COG,1808 |
| C35 C51 C73 C74 C77 C80 C84 C85 C120 C129 C130 | 2813189 | 2813189 | 2813189 | Capacitor, Cer, 10UF,+-20%,10V,X5R,0805 |
| C36 C59 C133 C134 C202 C203 C309 C315 C317 |  |  | 942529 | Capacitor,Ceramic,0.1uF,+-10%,25v,X7R,0805 |
| C37 C62 | 641453 | 641453 | 641453 | Capacitor, Cer, 15PF,+-10%,50V,C0G,0603 |
| C38 | 691516 | 691516 | 691516 | Capacitor, Cer,0.01uF,+-10%,500v,X7R,1210 |
| C39 C40 | 930235 | 930235 | 930235 | Capacitor, Cer,47pF,+-10%,1000v,COG,1808 |
| C4 | 855101 | 855101 | 855101 | Capacitor,Ceramic,22pF,+-5%,50v,COG,0805 |
| C44 C207 | 1619828 | 1619828 | 1619828 | Capacitor, Cer, 8.2PF,+-0.5PF,50V,C0G,0805 |
| C45 C46 C104 | 691508 | 691508 | 691508 | Capacitor,Ceramic, 100pF,+-10%,500v,COG,1206 |
| C48 C103 C307 |  |  | 740597 | Capacitor,Ceramic, 0.22uF,+80-20%,50v,Y5v,1206 |
| C49 C50 C222 C308 C310 C311 C314 C41 |  |  | 740597 | Capacitor,Ceramic,0.22uF,+80-20%,50v,Y5v,1206 |
| C61 C63 C66 C67 C81 C175 C223 | 644838 | 644838 | 644838 | Capacitor, Cer, 0.01UF,+-10%,50V,X7R,0603 |
| C64 C65 C68 C69 C71 C75 C76 C79 C173 C174 C199 C224 | 644838 | 644838 | 644838 | Capacitor, Cer, 0.01UF,+-10%,50V,X7R,0603,TAPE |
| C70 | 2387505 | 2387505 | 2387505 | Capacitor, Cer, 470PF,+-5%,50V,C0G,0603,TAPE |
| C78 | 2387505 | 2387505 | 2387505 | Capacitor, Cer, 470PF,+-5%,50V,C0G,0603,TAPE |
| C86 | 2092014 | 2092014 | 2092014 | Capacitor, Cer, 1UF,+-20%,16V,X5R,0603,TAPE |
| C87 C88 | 2151309 | 2151309 | 2151309 | Capacitor, Cer, 1UF,+-10%,10V,X5R,0603,TAPE |
| C101 C102 C182 C191 C192 | 855101 | 855101 | 855101 | Capacitor,Ceramic,22pF,+-5%,50v,COG,0805 |
| C106 | 106146 | 106146 | 106146 | Capacitor,Ceramic,0.01uF,+-10%,50v,X7R,0805 |
| C117 | 689460 | 689460 | 689460 | Capacitor, Cer, 0.22UF,+80-20%,16V,Y5V,0603,TAPE |
| C118 C119 | 1663963 | 1663963 | 1663963 | Capacitor, Tant, 47UF,+-20%,10V,3528,TAPE |
| C135 | 782615 | 782615 | 782615 | Capacitor, Cer, 0.047uF,+-20%,50v,X7R,1206 |
| C184-186 C193 | 1648909 | 1648909 | 1648909 | Capacitor, Cer, 0.22UF,+-10%,16V,X5R,0603,TAPE |
| C300 |  |  | 484378 | Capacitor, Cer, 1000pF,+-10%,50v,X7R,0805 |
| C301 |  |  | 943365 | Capacitor, Ceramic, 470pF,+-5%,50v,COG,1206 |
| C304 |  |  | 875914 | Capacitor, Cer,200pF,+-5%,50v,COG,0805 |
| C319 | 2131344 | 2131344 | 2131344 | Capacitor, Ceramic, 10 UF,+-10%,6.3v,X5R,0805 |
| CR2 CR105 CR107 CR110 CR300 CR301 |  |  | 742320 | Diode,Si,Pn,BAV99,70v,215ma,6ns,Dual,Series,SOT-23 |
| CR102 | 830489 | 830489 | 830489 | Diode,Si,BAS16, 75v,200ma,6ns,SOT-23 |
| CR103 CR104 | 107573 | 107573 | 107573 | Diode, Si, Pn, 400v,1a,1.8us,Do-214AC |
| CR108 CR109 | 929745 | 929745 | 929745 | Diode,Si,Schottky,Bat54s,30v,200ma,5ns,Dual-Series,SOT-23 |
| J1-4 | 943709 | 943709 | 943709 | Header, 2 Row, 2mm Ctr,20 Pin |
| J5 | 929229 | 929229 | 929229 | Connector, Flat Flex, 1mm Ctr, Rt Ang,16 Pos |
| K1 K2 K4 | 942917 | 942917 | 942917 | Relay, Armature, 2 Form C, 5Vdc,Latch |
| L1 L3-6 L8 L10 L13 L15 L16 | 601010 | 601010 | 601010 | Inductor, Bead, 600 Ohms@100mhz, 200mAdc, 1ohm, 1206 |
| L2 L7 L14 | 601010 | 601010 | 601010 | Inductor, Bead, 600 Ohms@100mhz, 200mAdc, 1ohm, 1206 |
| LS1 | 876995 | 876995 | 876995 | Audio Transducer, Piezo, Sounder, 2kHz, 10V, Washable, Trimmed Leads, 26x22mm |
| MP2 MP3 | 745802 | 745802 | 745802 | Fluke 52-8004, Contact, Thermocouple |
| Q1 Q5 Q9 Q12 Q106 | 742676 | 742676 | 742676 | Transistor, Npn, MMBT3904, 60V, 350mW, SOT-23 |
| Q2 Q7 Q8 Q10 Q11 Q13 Q107 | 742684 | 742684 | 742684 | Transistor, Si, Pnp, Mmbt3906, 40v, 200mA, 250mHz, 225mW, SOT-23 |
| Q3 Q4 | 930045 | 930045 | 930045 | Transistor, PNP, MUN2111, 50V, 200mW, SC-59 |
| Q6 | 1579628 | 1579628 | 1579628 | Mosfet, N,BSS123, 100V, 150MA, 6 Ohms, 250mW, SOT-23 |
| Q105 | 782490 | 782490 | 782490 | Transistor, N-Ch R05A, Si, N-Dmos, 500v, TO-92 |
| Q108 | 782508 | 782508 | 782508 | Mosfet, P-Chn R05A, Transistor,Si,P-Mos,500v,TO-92 |
| Q200 Q201 |  |  | 643681 | Transistor, Si, N-Mos, NDS351AN, 500mW, SOT-23 |
| R1 R159-161 | 804002 | 804002 | 804002 | Resistor,Cermet, 280,+-1%, 0.1W, 100PPM,0603 |
| R3 | 928825 | 928825 | 928825 | Resistor,Cermet,24.9k,+-1%,0.125W, 100ppm, 0805 |
| R4 R10 R26 R319 |  |  | 928713 | Resistor,Cermet,1k,+-1%,0.1W, 100ppm, 0805 |
| R7 R102 | 649720 | 649720 | 649720 | Resistor,Cermet, 1K,+-1%, 0.1W, 100PPM, 0603 |
| R8 R25 R52 R57 R65 R84 R85 R87 R91 R93 R99 | 605052 | 605052 | 605052 | Resistor,Cermet, 10K, +-1%, 0.1W, 100PPM, 0603 |
| R9 R56 R58 R64 | 2094195 | 2094195 | 2094195 | Resistor, MF, 200K,+-0.1%,0.1W, 50PPM, 0805 |
| R11 R236 | 928762 | 928762 | 928762 | Resistor, Cerm,4.02k,+-1%,0.1W, 100ppm, 0805 |
| R12 | 928924 | 928924 | 928924 | Resistor,Cermet,10,+-1%,0.125W, 100ppm, 0805 |
| R128 | 943642 | 943642 | 943642 | Resistor, Cerm,475,+-1%,0.1W, 100ppm, 0805 |
| R13 | 831305 | 831305 | 831305 | Resistor, Cerm,56.2k,+-1%,.125W, 100ppm, 1206 |
| R14 R109-111 R121 | 807172 | 807172 | 807172 | Resistor, Cerm,2k,+-1%,.125W, 100ppm, 1206 |
| R16 R17 R28 R33 R66 R68 R69 R72 R80 R116 R129 R147 R150 R219 R233 R239 R310 R311 | 928791 | 928791 | 928791 | Resistor,Cermet,10k,+-1%,0.125W, 100ppm, 0805 |
| R18 R32 R73 R74 R122-127 R315 |  |  | 928866 | Resistor,Cermet,100k,+-1%,0.125W, 100ppm, 0805 |
| R19 | 867085 | 867085 | 867085 | Resistor, Cerm,75k,+-1%,.125W, 100ppm, 1206 |
| R20 R21 | 928908 | 928908 | 928908 | Resistor, Cerm,66.5k,+-1%,0.1W, 100ppm, 0805 |
| R22 R101 | 605139 | 605139 | 605139 | Resistor,Cermet, 4.99K,+-1%,0.1W,100PPM,0603 |
| R23 | 1555198 | 1555198 | 1555198 | Resistor,Cermet, 20K,+-1%,0.1W,100PPM,0603 |
| R24 | 783274 | 783274 | 783274 | Resistor,Cermet,10m,+-5%,0.25W, 400ppm, 1206 |
| R30 R141 | 783274 | 783274 | 783274 | Resistor,Cermet,10m,+-5%,0.25W, 400ppm, 1206 |
| R31 | 642196 | 642196 | 642196 | Resistor,Cermet, 100,+-1%,0.1W,100PPM,0603 |
| R34 R53 | 689700 | 689700 | 689700 | Resistor,Cermet, 40.2K,+-1%,0.1W,100PPM,0603 |
| R35 | 643913 | 643913 | 643913 | Resistor,Cermet, 499K,+-1%,0.1W,100PPM,0603 |
| R36 R303 R306 |  |  | 944285 | Resistor, Cerm,499k,+-1%,0.1W, 100ppm, 0805 |
| R37 R41 | 928945 | 928945 | 928945 | Resistor,Cermet,1m,+-1%,0.125W, 100ppm, 0805 |
| R38 | 943311 | 943311 | 943311 | Resistor,Cermet,44.2k,+-1%,0.125W, 100ppm, 0805 |
| R39 R89 | 1711702 | 1711702 | 1711702 | Resistor,Cermet, 3.01K,+-1%,0.1W,100PPM,0603 |
| R40 R62 | 605052 | 605052 | 605052 | Resistor,Cermet, 10K,+-1%,0.1W,100PPM,0603 |
| R42 | 1562307 | 1562307 | 1562307 | Resistor,Cermet, 200K,+-1%,0.1W,100PPM,0603 |
| R43 | 691920 | 691920 | 691920 | Resistor, Mf,39.2k,+-1%,0.1W, 50ppm, 0805 |
| R44 | 649720 | 649720 | 649720 | Resistor,Cermet, 1K,+-1%,0.1W,100PPM,0603 |
| R45 | 2452646 | 2452646 | 2452646 | Resistor,Cermet, 10M,+-1%,0.1W,100PPM,0603 |
| R46 R55 | 867460 | 867460 | 867460 | Resistor, Cerm,665k,+-1%,.125W, 100ppm, 1206 |
| R47 | 687530 | 687530 | 687530 | Resistor,Metal Film,4.99k,\+-0.5%,0.125W, 10ppm, 1206 |
| R48 | 2151550 | 2151550 | 2151550 | Resistor,Cermet, 1.5M,+-1%,0.1W,100PPM,0603, |
| R49 R118 R15 R312 R317 R321 |  |  | 928866 | Resistor,Cermet,100k,+-1%,0.125W, 100ppm, 0805 |
| R50 | 643863 | 643863 | 643863 | Resistor,Cermet, 100K,+-1%,0.1W,100PPM,0603, |
| R51 | 811968 | 811968 | 811968 | Resistor,Cermet,15m,+-5%,0.25W, 300ppm, 1206 |
| R54 | 1562307 | 1562307 | 1562307 | Resistor,Cermet, 200K,+-1%,0.1W,100PPM,0603, |
| R59 | 2095352 | 2095352 | 2095352 | Resistor, MF, 4.99K,\+-0.1%,0.1W,50PPM,0805, |
| R60 | 650397 | 650397 | 650397 | Resistor, MF,100K,+-0.1%,0.1W,25PPM,0805 |
| R61 R77 | 944835 | 944835 | 944835 | Resistor, Var,Cerm,10k,+-25% |
| R63 | 810630 | 810630 | 810630 | Resistor, Cerm,1.5k,+-1%,.125W, 100ppm, 1206 |
| R70 R90 R95 R97 | 1555198 | 1555198 | 1555198 | Resistor,Cermet, 20K,+-1%,0.1W,100PPM,0603, |
| R71 R132 R134 | 1540992 | 1540992 | 1540992 | Resistor,Cermet, 10,+-1%,0.1W,100PPM,0603, |
| R75 R79 | 867692 | 867692 | 867692 | Resistor, Cerm,40.2k,+-1%,.125W, 100ppm, 1206 |
| R76 | 877006 | 877006 | 877006 | Resistor, Cerm,402k,+-1%,.125W, 100ppm, 1206 |
| R78 | 801258 | 801258 | 801258 | Resistor, Cerm,30.1k,+-1%,.125W, 100ppm, 1206 |
| R81 R234 | 928994 | 928994 | 928994 | Resistor,Cermet,4.75m,+-1%,0.125W, 400ppm, 0805 |
| R82 | 928093 | 928093 | 928093 | Resistor,Cermet, 50K,+-25% |
| R83 | 822312 | 822312 | 822312 | Resistor, Cerm,1.87k,+-1%,.125W, 100ppm, 1206 |
| R86 | 605193 | 605193 | 605193 | Resistor,Cermet, 11.3K,+-1%,0.1W,100PPM,0603, |
| R88 R96 | 605219 | 605219 | 605219 | Resistor,Cermet, 15.8K,+-1%,0.1W,100PPM,0603, |
| R92 | 803657 | 803657 | 803657 | Resistor,Cermet, 9.76K,+-1%,0.1W,100PPM,0603, |
| R94 | 605227 | 605227 | 605227 | Resistor,Cermet, 19.1K,+-1%,0.1W,100PPM,0603, |
| R98 | 1555180 | 1555180 | 1555180 | Resistor,Cermet, 7.5K,+-1%,0.1W,100PPM,0603, |
| R100 R130 R135 R67 | 1578104 | 1578104 | 1578104 | Resistor,Cermet, 5.11,+-1%,0.1W,400PPM,0603, |
| R103 R104 R235 | 867697 | 867697 | 867697 | Resistor, Cerm,150k,+-1%,.125W, 100ppm, 1206 |
| R105 | 807172 | 807172 | 807172 | Resistor, Cerm,2k,+-1%,.125W, 100ppm, 1206 |
| R106 R156 R162 R223 | 1598173 | 1598173 | 1598173 | Resistor,Cermet, 4.02K,+-1%,0.1W,100PPM,0603, |
| R107 R108 R113 R115 R2 R220 R29 R308 R5 R6 |  |  | 928791 | Resistor,Cermet,10k,+-1%,0.125W, 100ppm, 0805 |
| R112 | 928713 | 928713 | 928713 | Resistor,Cermet,1k,+-1%,0.1W, 100ppm, 0805 |
| R114 R117 | 928945 | 928945 | 928945 | Resistor,Cermet,1m,+-1%,0.125W, 100ppm, 0805 |
| R119 | 928882 | 928882 | 928882 | Resistor,Cermet,200k,+-1%,0.1W, 100ppm, 0805 |
| R131 | 604394 | 604394 | 604394 | Resistor,Cermet, 0,+0.05 Jumper, MAX,0.1W,0603, |
| R133 | 1540992 | 1540992 | 1540992 | Resistor,Cermet, 10,+-1%,0.1W,100PPM,0603, |
| R136 R137 | 943654 | 943654 | 943654 | Resistor,Cermet, 1,+-5%,0.1W,400PPM,0603, |
| R138 | 604956 | 604956 | 604956 | Resistor,Cermet, 10K,+-5%,0.1W,200PPM,0603, |
| R145 R146 R152 | 876524 | 876524 | 876524 | Resistor, Cerm,226k,+-1%,.125W, 100ppm, 1206 |
| R148 R151 R27 | 807735 | 807735 | 807735 | Resistor, Cerm,27,+-5%,.125W, 200ppm, 1206 |
| R154 | 688838 | 688838 | 688838 | Resistor,Bulk Metal Foil,200,+-1%,0.25W, 3ppm,6032 |
| R155 | 928825 | 928825 | 928825 | Resistor,Cermet,24.9k,+-1%,0.125W, 100ppm, 0805 |
| R166 | 2151550 | 2151550 | 2151550 | Resistor,Cermet, 1.5M,+-1%,0.1W,100PPM,0603, |
| R180 R226-232 R313 |  |  | 927707 | Resistor, Cerm,47,+-5%,.0625W, 200ppm,0603 |
| R224 R225 | 928762 | 928762 | 928762 | Resistor, Cerm,4.02k,+-1%,0.1W, 100ppm, 0805 |
| R238 | 746347 | 746347 | 746347 | Resistor, Cerm,220,+-5%,.125W, 200ppm, 1206 |
| R300 |  |  | 877006 | Resistor, Cerm,402k,+-1%,.125W, 100ppm, 1206 |
| R301 R305 |  |  | 866830 | Resistor, Cerm,115k,+-1%,.125W, 100ppm, 1206 |
| R302 |  |  | 887112 | Resistor, Cerm,487k,+-1%,.125W, 100ppm, 1206 |
| R304 R314 |  |  | 836643 | Resistor, Cerm,215k,+-1%,.125W, 100ppm, 1206 |
| R307 |  |  | 836437 | Resistor, Cerm,3m,+-5%,.125W, 200ppm, 1206 |
| R309 |  |  | 928924 | Resistor,Cermet,10,+-1%,0.125W, 100ppm, 0805 |
| R316 |  |  | 746644 | Resistor, Cerm,20k,+-5%,.125W, 200ppm, 1206 |
| R318 |  |  | 746768 | Resistor, Cerm,300k,+-5%,.125W, 200ppm, 1206 |
| R320 |  |  | 816033 | Resistor, Cerm,14k,+-1%,.125W, 100ppm, 1206 |
| R322 |  |  | 944439 | Resistor,Cermet,0.22,+-5%,0.125W, 600ppm, 0805 |
| R323 | 867432 | 867432 | 867432 | Resistor, Cerm,562k,+-1%,.125W, 100ppm, 1206 |
| U1 | 867932 | 867932 | 867932 | IC, Op Amp, Dual, LM358, Low Power, Soic |
| U2 | 929315 | 929315 | 929315 | IC, Op Amp, TC913, 30uV Offset, Dual, Chopper-Stab, SO8 |
| U3 | 913835 | 913835 | 913835 | IC, Cmos, Comparator TLC393, Dual, uPower, SO8 |
| U4 | 929828 | 929828 | 929828 | IC, Op Amp, AD820A, 5V To 36V, 0.8mV Offset, 1.8mHz, R/R, S/S, Lo Pwr, SO8 |
| U6 | 2570483 | 2570483 | 2570483 | IC, Programmable Waveform Generator,AD9833, 2.3-5.5V, Msop10 |
| U7 | 107565 | 107565 | 107565 | IC, Cmos, LTC1150 Op Amp, Chopper, +/-15v, Int Caps, SO8 |
| U8 U9 U101 U102 |  | 914101 | 914101 | IC, Memory, Sram, K6T1008, 1MB, 128kx8, 5v, 100ns, SO32 |
| U10-12 U5 | 876503 | 876503 | 876503 | IC, Cmos LTC1050 Op Amp, Chopper, Int Caps, SO8 |
| U13-16 | 876503 | 876503 | 876503 | IC, Cmos LTC1050 Op Amp, Chopper, Int Caps, SO8 |
| U22 | 2437589 | 2437589 | 2437589 | IC,Analog Switch, ADG619,2.7-5.5V,6 Ohms,SPDT,190MHz,SOT-23-8, |
| U23 | 3086038 | 3086038 | 3086038 | IC, Analog Switch, ADG1208, 200 Ohms, 8:1 MUX/DEMUX, TSSOP16, |
| U24 | 3087267 | 3087267 | 3087267 | IC,Microcontroller,18F2410,8 BIT,5V,40MHZ,16K FLASH,QFN28, |
| U30 | 1610921 | 1610921 | 1610921 | IC, Analog Switch, 7SB3157,1.65-5.5V,7 Ohms, SPDT, SC70-6 |
| U31 | 1990272 | 1990272 | 1990272 | IC,Op Amp, AD623, 500uV Offset, 800KHZ, USO8 |
| U32 | 643368 | 643368 | 643368 | IC, Op Amp, LMC6042, 3mV Offset, 100KHz, Cmos, Dual,SO8 |
| U33 U39 | 1605502 | 1605502 | 1605502 | IC, Op Amp, AD8541, 6mV Offset, 1MHz, SOT-23-5, |
| U34 | 2554038 | 2554038 | 2554038 | IC, Analog Switch, ADG659, 100 Ohms, Dual 4:1 MUX/DEMUX, TSSOP16 |
| U35 | 2113880 | 2113880 | 2113880 | IC, Op Amp, LT1880, 200uV Offset, 1.1MHz, SOT-23-5, |
| U36 | 3086212 | 3086212 | 3086212 | IC, Op Amp, LT6011A,2.7TO+-18V, 60uV Offset, 330KHZ, Dual, 400nVP-P Noise |
| U38 | 1573086 | 1573086 | 1573086 | IC,Voltage Regulator, Linear, LT1761,ADJ, 100MA, LDO, Low Power, SOT-23-5, |
| U40 | 2151412 | 2151412 | 2151412 | IC,ADC,LTC2440-1,5V, 24-BIT, Diff, Serial, Delta-Sigma,SSOP16, |
| U41 | 3271828 | 3271828 | 3271828 | IC,Voltage Reference, LT1021, 5V, 1%, 5PPM/C, 1.2mA,TO-5 |
| U42-44 | 690765 | 690765 | 690765 | IC, Cmos 7SZ125 Single Buffer W/3-ST,SOT-23-5 |
| U45 | 1563433 | 1563433 | 1563433 | IC,Analog Switch,DG413,3-44V, 35 Ohms, SPST, Quad, 2NC, 2NO, SO16, |
| U46 | 643616 | 643616 | 643616 | IC,Cmos, 6061 Op Amp, Precision, uPower, SOIC8 |
| U47 | 3157659 | 3157659 | 3157659 | IC, Op Amp, LTC6240HV, 60uV Offset, 1pA,18MHz, SO8, |
| U100 | 929083 | 929083 | 929083 | IC, Microprocessor, 68306, 16 Bit, 5v, 16mHz, Integrated, PQFP132 |
| U103 U104 | 107562 | 107562 | 688827 | IC, Memory, Flash, 28F008BV-B, 8MB, 1Mx8, 5V, 80ns, Bottom Boot, TSOP40 |
| U106 | 3100445 | 3100445 | 3100445 | IC,Real Time Clock-Calendar, 72423A, 20PPM, SO24 |
| U107 | 3248778 | 3248778 | 3248778 | IC,LCD Driver,RA8835A, Graphic Lcd Controller, TQFP60 |
| U108 | 929609 | 929609 | 929609 | IC, Memory, Sram, 43256, 256kb, 32kx8, 5v, 70ns, SO28 |
| U109 U110 | 852090 | 852090 | 852090 | IC, Cmos 74HC126 Quad Buffer W/3-St, Soic |
| U111 U112 | 852090 | 852090 | 852090 | IC, Cmos 74HC126 Quad Buffer W/3-St, Soic |
| U114 | 929224 | 929224 | 929224 | IC, Cmos 694 Microprocessor Supervisor, SO8 |
| U115 | 866785 | 866785 | 866785 | IC, Cmos, 68C681,Dual Channel Uart, Plcc |
| U117 | 867932 | 867932 | 867932 | IC, Op Amp, Dual, LM358, Low Power, Soic |
| U118 | 807040 | 807040 | 807040 | IC, Bipolar 1403 2.5v,40 Ppm T.C.,Bandgap Ref, Soic |
| U119 | 929273 | 929273 | 929273 | IC Cmos 74HC4094 8bit Shft Rgs W/3-St, Soic |
| U120 | 3317608 | 3317608 | 3317608 | IC,LOGIC,74HC112,DUAL JK F/F,-EDG TRG,TSSOP16, |
| U121 | 929299 | 929299 | 929299 | IC, Cmos, MAX529 8 Bit Ser Dac,1 LSB, Soic |
| U123 | 876813 | 876813 | 876813 | IC, Cmos 74ACT14 Hex Inverter W/Scht Trig, Soic |
| U124 | 929075 | 929075 | 929075 | IC, Op Amp 34072 Dual, High Bw, Sngl Sup, SO8 |
| U125 | 642436 | 642436 | 642436 | IC, Analog Switch, ADG433, 12 To -15V, 24 Ohms, Spst, Quad, No, Isolated, So16 |
| U126 | 107524 | 107524 | 107524 | IC, Cmos ADG513 Quad Bilateral Switch, Soic |
| U128 | 2132280 | 2132280 | 2132280 | IC, Dac, LTC1590, 5.0v,12-Bit, Dual, Current, Serial, Soic16 |
| U129 U170 |  |  | 806794 | IC,Op Amp TI062 Dual, Jfet, Low Power, Soic |
| U130 U134 | 107524 | 107524 | 107524 | IC, Cmos ADG513 Quad Bilateral Switch, Soic |
| U135 | 942701 | 942701 | 942701 | IC, Cmos, ADG511 Quad Bilateral Switch,Soic |
| U137 | 2089072 | 2089072 | 2089072 | IC,RMS Converter, AD737J-5,+-2.5V, SO8 |
| U141 U19 U28 U29 | 942680 | 942680 | 942680 | Isolator, Opto, HCPL:-0501, LED To Transistor, Single Channel, SO8 |
| U163 | 942698 | 942698 | 942698 | IC,Voltage, Bipolar 1021 Ref,7 V,+-50 Mv,10 Ppm, SO8 |
| U171 | 837054 | 837054 | 837054 | IC, Logic, 74HC393, 2-6V, Divide By 16 Binary Counter, Dual, So14 |
| U172 |  |  | 2061960 | IC, Comm, A5191HRT, Hart Modem, 3.3V-5.0V, Plcc28 |
| U173 U175 | 866801 | 866801 | 866801 | IC, Cmos 74HC125 Quad Bus Buffer W/3-St, Soic |
| U174 |  |  | 929489 | IC,Voltage Reference, LM4041, 1.2V, 2%, 150ppm/C, 0.065-20mA, SOT-23 |
| VR3 VR4 | 837187 | 837187 | 837187 | Zener, Uncomp, MMBZ5245B, 15W, 5%, 8.5mA, 225mW, SOT-23 |
| VR104 VR105 | 837161 | 837161 | 837161 | Zener, Uncomp, MMBZ5233B, 6V, 5%, 20mA, 225mA, SOT-23 |
| WP1 WP2 | 2563832 | 2563832 | 2563832 | Standoff, 4-40, .187, Steel, Tin, Smd |
| XBT1 | 929448 | 929448 | 929448 | Holder, Battery, Pwb, 2032 Coin |
| Y1 | 929245 | 929245 | 929245 | Crystal, 3.84mHz, 50/100ppm, 20pF, Plastic Encapsulated, Smd |
| Y2 | 929240 | 929240 | 929240 | Crystal, 3.6864mHz, 100/100ppm, 20pF, Plastic Encapsulated, Smd |
| Y3 | 943105 | 943105 | 943105 | Crystal,10mhz,50/100ppm,20pF,Plastic Encapsulated,Smd |
| Z1 | 655811 | 655811 | 655811 | RNET,MF,Poly, SIP,8840 LO V Divider |
| Z4 | 3277485 | 3277485 | 3277485 | RNET,MF, Poly, DIP, HI V Inst Amp |
| Z5 | 3272109 | 3272109 | 3272109 | RNET, MF, Poly, SIP, LO V Divider |
| Z6 | 103143 | 103143 | 103143 | Rnet, Mf, Poly, Dip, Lo V Inst Amp |
| Z11 Z12 | 626462 | 626462 | 626462 | Rnet, Thin Film, Tested, Smr |

Power Supply PCA Assembly

A few parts changed on the Power Supply PCA for the Second Version of the Main PCA.  Parts used when paired with the  first version  Main PCA are shown in Red in this parts list.

| Reference Designator | Item Number | Power Supply Component Description |
| --- | --- | --- |
| BC1 BC3 BC5 | 938303 | Contact, PCB Battery Interface |
| C1 C28 C55 | 876545 | Capacitor, Tant, 22UF,+-20%,6V,6032 |
| C2 C20 C25-26 C50 C58 C64 C83 | 740597 | Capacitor, Cer, 0.22UF,+80-20%,50V,Y5V,1206 |
| C3 C10-11 | 927814 | Capacitor, Tant, 10UF,+-20%,25V,6032 |
| C4 C8 C12-13 C23 C32 C52 C82 | 740597 | Capacitor, Cer, 0.22UF,+80-20%,50V,Y5V,1206 |
| C5 C46 | 943589 | Capacitor, Tant, 4.7UF,+-20%,50V,7343 |
| C6 | 691516 | Capacitor, Cer, 0.01UF,+-10%,500V,X7R,1210 |
| C7 C81 C84 | 691417 | Capacitor, Cer, 1000PF,+-10%,500V,X7R,1206 |
| C9 | 769786 | Capacitor, Cer, 1800PF,+-10%,50V,C0G,1206 |
| C14-17 C34-35 C39 C42 C48 | 876552 | Capacitor, Tant, 3.3UF,+-20%,16V,3528 |
| C18-19 C24 C61 C80 | 867650 | Capacitor, Cer, 100PF,+-1%,50V,C0G,0805 |
| C21 | 758078 | Capacitor, Cer, 220PF,+-10%,50V,C0G,1206 |
| C22 C47 | 876545 | Capacitor, Tant, 22UF,+-20%,6V,6032 |
| C27 | 927814 | Capacitor, Tant, 10UF,+-20%,25V,6032 |
| C29 | 691417 | Capacitor, Cer, 1000PF,+-10%,500V,X7R,1206 |
| C30 C41 C45 C57 C59 C63 C65-67 C69-71 C73-76 C86-87 | 942529 | Capacitor, Cer, 0.1UF,+-10%,25V,X7R,0805 |
| C31 | 769240 | Capacitor, Cer, 33PF,+-10%,50V,C0G,1206 |
| C33 C43 C56 C62 C68 C72 C90 | 942529 | Capacitor, Cer, 0.1UF,+-10%,25V,X7R,0805 |
| C36-38 C40 | 876552 | Capacitor, Tant, 3.3UF,+-20%,16V,3528 |
| C44 | 811687 | Capacitor, Cer, 0.027UF,+-5%,50V,X7R,1206 |
| C49 | 484378 | Capacitor, Cer, 1000PF,+-10%,50V,X7R,0805 |
| C51 | 106146 | Capacitor, Cer, 0.01UF,+-10%,50V,X7R,0805 |
| C53-54 | 769240 | Capacitor, Cer, 33PF,+-10%,50V,C0G,1206 |
| C60 | 484378 | Capacitor, Cer, 1000PF,+-10%,50V,X7R,0805 |
| C77-78 | 867650 | Capacitor, Cer, 100PF,+-1%,50V,C0G,0805 |
| C79 C85 | 644960 | Capacitor, Cer, 0.33UF,+-5%,25V,X7R,1206 |
| CR1 | 944264 | Diode, MURS120 SMR,Diode, SI,ULTRAFAST,200V,1A,SMB |
| CR2-5 CR12 CR20 | 605821 | Diode, MBRS140 SMR,Diode, SI,SCHOTTKY,40V,1A,DO-214AB |
| CR6-8 CR11 CR13 CR16 | 742320 | Diode, SI,PN,BAV99,70V,215MA,6NS,DUAL,SERIES,SOT-23 |
| CR9-10 CR15 | 830489 | Diode, SI,BAS16,75V,200MA,6NS,SOT-23 |
| CR14 | 867072 | Diode, MMBD1501 SMR,Diode, SI,150V,200MA,SOT-23 |
| CR17 CR19 | 107573 | Diode, SI,PN,S1G,400V,1A,1.8US,DO-214AC |
| CR18 | 107573 | Diode, SI,PN,S1G,400V,1A,1.8US,DO-214AC |
| J8 | 688819 | Header, 2 Row, 2mm Ctr, 26 Pin |
| K1 | 942917 | Relay, Armature, 2 Form C, 5Vdc, Latch |
| L1 L3 | 929211 | Inductor,100UH,+-20%,0.51ADC |
| P2-4 | 2726686 | Cable Assembly,Flat,2 Row 2mm ctr Molded Header/Socket,20 Cond,1.20 L |
| Q1 Q5 Q8 | 642089 | MOSFET,SI,N,IRFR320,400V,3.1A,1.8 Ohm, 2.5W,D-PAK |
| Q2 | 836544 | FET,IRFR9020 SMR,Transistor, P-MOS,1W,D-PAK |
| Q3 Q10-11 | 643681 | FET,N,NDS351AN SMR,Transistor, N-MOS,500MW,SOT-23 |
| Q4 Q6-7 Q9 Q12-13 Q18 Q26 | 742676 | NPN,MMBT3904 SMR,Transistor, NPN,60V,350MW,SOT-23 |
| Q14 Q16 | 802626 | MOSFET,SI,N,DN2540,400V,170MA,25 Ohm, 1.6W,DEPLETION MODE,SOT-89 |
| Q15 Q19-20 | 742684 | Transistor, PNP,MMBT3906,40V,200MA,250MHZ,225MW,SOT-23 |
| Q17 | 3360660 | Transistor, NPN,CXTA44,450V,300MA,20MHZ,1.2W,SOT-89 |
| Q17 | 603289 | Transistor, BF620, NPN, 300V, 1W, SOT-89 (1st Vers) |
| Q21 | 3360613 | Transistor, PNP,FCX558,400V,200MA,50MHZ,1W,SOT-89 |
| Q21 | 603297 | Transistor, PNP,BF621 300V, 1W, SOT-89 (1st Vers) |
| Q22 Q24 | 643681 | FET,N,NDS351AN Transistor, N-MOS,500MW,SOT-23 |
| Q23 | 742676 | NPN,MMBT3904 Transistor, NPN,60V,350MW,SOT-23 |
| Q25 | 876263 | N-JFET Transistor, N-JFET,SEL,SOT-23 |
| R1 | 928937 | Resistor, Cermet, 100,+-1%,0.125W,100PPM,0805 |
| R2 | 877006 | Resistor, Cermet, 402K,+-1%,.125W,100PPM,1206 |
| R3 R7 R51 | 887109 | Resistor, Cermet, 43.2K,+-1%,0.25W,100PPM,1206 |
| R4 | 911243 | Resistor, Cermet, 324K,+-1%,.125W,100PPM,1206 |
| R5 R26 R44 | 928945 | Resistor, Cermet, 1M,+-1%,0.125W,100PPM,0805 |
| R8 R53 | 928796 | Resistor, Cermet, 11K,+-1%,0.1W,100PPM,0805 |
| R9 R18 R62 | 928762 | Resistor, Cermet, 4.02K,+-1%,0.1W,100PPM,0805 |
| R10 R35 | 746347 | Resistor, Cermet, 220,+-5%,.125W,200PPM,1206 |
| R11 R14 R46-47 | 867692 | Resistor, Cermet, 40.2K,+-1%,.125W,100PPM,1206 |
| R12 R50 R55 R70 R77 R103 | 928866 | Resistor, Cermet, 100K,+-1%,0.125W,100PPM,0805 |
| R13 R54 R66-67 R75 R78 | 928791 | Resistor, Cermet, 10K,+-1%,0.125W,100PPM,0805 |
| R15 | 447011 | Resistor, MF,16.5,+-1%,0.125W,50PPM |
| R16 | 929310 | Resistor, MF,2.7,+-5%,0.125W,200PPM |
| R17 | 930110 | Resistor, Cermet, 5.62K,+-1%,0.125W,100PPM,0805 |
| R19 R68 | 887109 | Resistor, Cermet, 43.2K,+-1%,0.25W,100PPM,1206 |
| R21-22 R93 | 928825 | Resistor, Cermet, 24.9K,+-1%,0.125W,100PPM,0805 |
| R23-24 | 811885 | Resistor, Cermet, 432,+-1%,0.25W,100PPM,1206 |
| R25 | 867697 | Resistor, Cermet, 150K,+-1%,.125W,100PPM,1206 |
| R27 | 866830 | Resistor, Cermet, 115K,+-1%,.125W,100PPM,1206 |
| R28 | 2095352 | Resistor, MF, 4.99K,\+-0.1%,0.1W,50PPM,0805 |
| R28 | 943030 | Resistor, Cermet, 5.76K,+-1%,0.1W,100PPM,0805 (1st Vers) |
| R29 | 867692 | Resistor, Cermet, 40.2K,+-1%,.125W,100PPM,1206 |
| R30 R34 | 867085 | Resistor, Cermet, 75K,+-1%,.125W,100PPM,1206 |
| R31 | 811885 | Resistor, Cermet, 432,+-1%,0.25W,100PPM,1206 |
| R32 | 876524 | Resistor, Cermet, 226K,+-1%,.125W,100PPM,1206 |
| R33 | 867460 | Resistor, Cermet, 665K,+-1%,.125W,100PPM,1206 |
| R36 | 810788 | Resistor, Cermet, 3.32K,+-1%,.125W,100PPM,1206 |
| R38 | 928924 | Resistor, Cermet, 10,+-1%,0.125W,100PPM,0805 |
| R40 R89 | 867085 | Resistor, Cermet, 75K,+-1%,.125W,100PPM,1206 |
| R41-43 | 928713 | Resistor, Cermet, 1K,+-1%,0.1W,100PPM,0805 |
| R45 R56 R71 R73 R76 R80 R83 R88 R91 | 876524 | Resistor, Cermet, 226K,+-1%,.125W,100PPM,1206 |
| R48 | 687530 | Resistor, MF, 4.99K,\+-0.5%,0.125W,10PPM,1206 |
| R48 | 801258 | Resistor, Cermet, 30.1K,+-1%,.125W,100PPM,1206 (1st Vers) |
| R49 | 928697 | Resistor, Cermet, 49.9K,+-1%,0.1W,100PPM,0805 |
| R52 | 943316 | Resistor, Cermet, 15K,+-1%,0.1W,100PPM,0805 |
| R58 R60 R90 R92 R97 R102 | 928713 | Resistor, Cermet, 1K,+-1%,0.1W,100PPM,0805 |
| R59 | 928705 | Resistor, Cermet, JUMPER,0,+0.05 MAX,0.125W,0805 |
| R59 | 943071 | Resistor, Cermet, 24.3K,+-1%,0.1W,100PPM,0805 (1st Vers) |
| R6 R20 R37 R39 R57 R69 R79 R82 | 928791 | Resistor, Cermet, 10K,+-1%,0.125W,100PPM,0805 |
| R61 R87 | 928866 | Resistor, Cermet, 100K,+-1%,0.125W,100PPM,0805 |
| R63 | 928937 | Resistor, Cermet, 100,+-1%,0.125W,100PPM,0805 |
| R64 | 783274 | Resistor, Cermet, 10M,+-5%,0.25W,400PPM,1206 |
| R65 | 801258 | Resistor, Cermet, 30.1K,+-1%,.125W,100PPM,1206 |
| R72 | 928825 | Resistor, Cermet, 24.9K,+-1%,0.125W,100PPM,0805 |
| R74 | 928861 | Resistor, Cermet, 61.9K,+-1%,0.1W,100PPM,0805 |
| R81 | 944038 | Resistor, MF, 10,+-1%,0.25W,6PPM,6032 |
| R84 | 944970 | Resistor, Wirewound, 1K,+-1%,1.5W,20PPM,6538 |
| R85 | 944285 | Resistor, Cermet, 499K,+-1%,0.1W,100PPM,0805 |
| R86 | 866830 | Resistor, Cermet, 115K,+-1%,.125W,100PPM,1206 |
| R94 | 602711 | Resistor, Cermet, 301K,+-1%,0.1W,100PPM,0805 |
| R95-96 | 928945 | Resistor, Cermet, 1M,+-1%,0.125W,100PPM,0805 |
| R98-100 | 641156 | Resistor, Cermet, 698,+-1%,0.1W,100PPM,0805 |
| T1 | 665133 | Transformer Switching |
| U1 | 929190 | IC,Volt Reg, 2936, 5V, 50mA, SO8 |
| U2 U6 | 944249 | IC,Volt Reg, 1172, Switching, 100KHz, 1.2A, SO8 |
| U3 U7 | 867932 | IC,Op Amp, LM358, Dual, SOIC |
| U4 | 782995 | IC,Cmos, 74HC74, Dual D F/F,+Edge Trg,SOIC |
| U5 U9 | 942693 | IC,VOLT REG, 2937, 5V, 500mA, TO-263 |
| U8 | 929075 | IC,Op Amp, 34072, Dual, High Bw, Sngl Sup, SO8 |
| U10 | 944996 | IC,Volt Reg, LP2980, 5V, 50mA, SOT-23-5 |
| U11 U23-24 | 107524 | IC,Cmos, ADG513 Quad Bilateral Switch, SOIC |
| U12 | 2132280 | IC,DAC, LTC1590, 5.0V, 12-Bit, Dual, Current, Serial, SOIC16 |
| U13 | 929273 | IC,Cmos, 74HC4094 8 Bit Shft Rgs W/3-St, SOIC |
| U14-15 | 929281 | IC,Opto-Isolator,LED To Transistor,SO4 |
| U16 U18 | 107554 | IC,Cmos, 7211U50 VOLT REG, -5.0 V, 40 MA, SOT-89 |
| U17 | 942680 | Opto-Isolator, HCPL-0501,LED To Transistor, Single Channel,SO8 |
| U19 | 942698 | IC,Voltage Ref, 1021 7 V, +-50 MV, 10 PPM, SO8 |
| U20-21 | 929315 | IC,Op Amp, TC913, 30 uV Offset, 1.5MHZ, DUAL, CHOPPER-STAB, SO8 |
| U25 | 838177 | IC,Cmos, 74ACT00, Quad 2 Input Nand Gate, SOIC |
| U27 | 876503 | IC,Cmos, LTC1050, Op Amp,Chopper, Int Caps, SO8 |
| U28 | 929828 | IC,OP AMP,AD820A, 0.8mV Offset, 1.8MHz, SO8 |
| VR1 | 831230 | Zener, MMBZ5251B, 22V,5%,5.6mA,225mW,SOT-23 |
| VR2 | 875604 | Zener, MMBZ5232B, 5.6V, 5%, 20mA, 225mW, SOT-23 |
| VR3 | 943456 | Zener, MMBZ5258B, 36V, 5%, 3.4mA, 225mW, SOT-23 |
| VR5 | 837146 | Zener, MMBZ5237B, 8.2V, 5%, 20mA, 225mW, SOT-23 |
| VR6 | 943456 | Zener, MMBZ5258B, 36V, 5%, 3.4mA, 225mW, SOT-23 |
| VR7-8 | 837161 | Zener, MMBZ5233B, 6V, 5%, 20mA, 225mW, SOT-23 |
| Z1 | 645341 | RNET, MF, POLY, SIP, 1752 Low V Divider |
| Z2 | 655811 | RNET, MF, POLY, SIP, 8840 Low V Divider |

I/O PCA Assembly

| Ref. Des. | 741B PN | 743B PN | 744 2020 PN | Description |
| --- | --- | --- | --- | --- |
| C1 | 605425 | 605425 | 605425 | Capacitor, Tant, 2.2UF,+-20%,10V,3216 |
| C2-5 C9 | 942529 | 942529 | 942529 | Capacitor, Ceramic, 0.1UF,+-10%,25V,X7R,0805 |
| C6 | 605425 | 605425 | 605425 | Capacitor, Tant, 2.2UF,+-20%,10V,3216 |
| C7-8 C10 | 942529 | 942529 | 942529 | Capacitor, Ceramic, 0.1UF,+-10%,25V,X7R,0805 |
| H007-008 | 558791 |  |  | Screw, 4-40, 0.250, Pan, Phillips, Stainless, Blk Oxide |
| H1-2 | 2583053 | 2583053 | 2583053 | Standoff, 4-40, 0.375, Steel, Tin, Smd, Tape |
| H19-20 |  | 448092 | 448092 | Connector Accessory Kit,D-Sub Jack Screw,4-40 |
| J7 | 942730 | 942730 | 942730 | Jack,Panel/Pwb,2.54 Mm,Pin |
| L1 L3 L5 | 601069 | 601069 | 601069 | Inductor, CM, 1500 Ohms @100 Mhz,2020 |
| L2 | 944645 | 944645 | 944645 | Inductor,Bead,60 Ohms@100mhz,6adc,10mohm,1806 |
| L4 | 601010 | 601010 | 601010 | Inductor,Bead,600 Ohms@100mhz,200madc,1ohm,1206 |
| L6-7 |  |  | 601010 | Inductor,Bead,600 Ohms@100mhz,200madc,1ohm,1206 |
| MP2 | 938253 | 938253 | 938253 | 7xx I/O Bracket |
| MP3 | 938266 |  |  | 7xx I/O Plug (Cover) |
| P3 | 942714 | 942714 | 942714 | Connector, Circular, Panel, 5 Socket |
| P4 |  | 942727 | 942727 | Connector, D-SUB, 9 Socket |
| R1 R4 R7 R10 | 943642 | 943642 | 943642 | Resistor, Cermet, 475,+-1%, 0.1W, 100PPM, 0805 |
| R2-3 R5 R9 | 928762 | 928762 | 928762 | Resistor, Cermet, 4.02K,+-1%, 0.1W, 100PPM, 0805 |
| R6 | 928762 | 928762 | 928762 | Resistor, Cermet, 4.02K,+-1%, 0.1W, 100PPM, 0805 |
| R8 | 943311 | 943311 | 943311 | Resistor, Cermet, 44.2K,+-1%, 0.125W, 100PPM, 0805 |
| R11 | 943613 | 943613 | 943613 | Resistor, Cermet, 1.47K,+-1%, 0.1W, 100PPM, 0805 |
| R12 | 943613 | 943613 | 943613 | Resistor, Cermet, 1.47K,+-1%, 0.1W, 100PPM, 0805 |
| R13-14 | 943006 | 943006 | 943006 | Resistor, Cermet, 2.1K,+-1%, 0.1W, 100PPM, 0805 |
| T1 |  |  | 662848 | Transformer, Coupling,1:1,1.6H,FLUKE-744-6301,DIP10 |
| U1-2 U4-5 | 942680 | 942680 | 942680 | Isolator, Opto, HCPL-0501,LED To Transistor,Single Channel,SO8 |
| U3 | 929281 | 929281 | 929281 | Isolator, Opto, LED To Transistor,SO4 |
| U6 | 821538 | 821538 | 821538 | IC, Cmos 145406, RS232 Driver/Receiver,SOIC |
| U7 | 944996 | 944996 | 944996 | IC, Volt Reg LP2980, 5V, 50mA, SOT-23-5 |
| U8-9 | 929286 | 929286 | 929286 | Isolator, Opto, LED To Darlington, SOIC |
| U10 | 691261 | 691261 | 691261 | IC, Volt Reg LT1175, FIXED,-5V, 500mA, SO8 |
| U11 | 944996 | 944996 | 944996 | IC, Voltage Regulator,Linear,Lp2980, 50mA, SOT-23-5 |

Thermal PCA Assembly

| Ref. Des. | Item Number | Description |
| --- | --- | --- |
| C1 | 106146 | Capacitor, Ceramic, 0.01uF,+-10%,50V |
| E1 E2 | 938204 | 74x T/C Contact |
| L1 L2 | 601010 | Ferrite Chip,600 Ohm @100 MHz |
| P2 | 943464 | Header,1 Row, .100ctr, Rt Ang, 4 Pin |
| R1 | 604311 | Resistor, Cermet, 20,+-1%,.125W,100PPM,1206 |
| RT2 | 942685 | Thermistor, Chip, Neg, 10K,+-10%,25C |